## Slide 1: Tales from the Field:
Agile in the Regulatory Context
Copyright © 2016 by CMD MedTech, LLC. Permission granted to INCOSE to publish and use.

## Slide 2: Agenda
Background
Regulations, Standards, and Guidance
Experiences – What Works Well?
  Pre-Market
  Post-Market
  Auditor / Inspector Comments
Questions and Discussion

## Slide 3: Quick Background
About Me
About CMD MedTech

## Slide 5: Regulatory Overview
  FDA is focus of this presentation
  Software-centric topics, but also cover systems-level agile practices
  FDA generally indicates “whats” not “hows”
  FDA participated in and recognizes TIR45:2012 - Guidance on the use of agile practices in the development of medical device software

## Slide 6: Regulation
Most all sections of 21 CFR 820.30 (Design Controls) are touched by agile development:
  Design and Development Planning
  Design Input
  Design Output
  Design Review
  Design Verification
  Design Validation
  Design Changes
  Design History File
Review the Quality System Regulation preamble! States FDA’s intent on QSR development and subsequent enforcement
  The word “flexibility” is used 56 times.
QSIT: Quality System Inspection Technique. Used by FDA to guide inspections

## Slide 7: Key Guidance
FDA Software:
  Guidance for the Content of Software in Premarket Submissions
  General Principles of Software Validation
  Mobile Medical Applications
  Cybersecurity in Premarket Submissions
  Off-the-Shelf Software
  Human Factors
  General Wellness: Policy for Low Risk Devices
Other FDA:
  When to re-submit a 510(k)
AAMI
  AAMI TIR45:2012 Guidance on the use of agile practices in the development of medical device software

## Slide 8: Key Standards
EN / IEC 62304: Medical Device Software – Software Lifecycle Process
  The primary standard for medical device software development
EN / IEC 60601-1: Medical electrical equipment - Part 1: General requirements for basic safety and essential performance
  References IEC 62304, and contains additional requirements
EN / ISO 14971: Medical devices -- Application of risk management to medical devices
EN / IEC 62366: Medical devices -- Application of usability engineering to medical devices

## Slide 9: Regulations, Standards, Guidance
Following regulation and applicable guidance and standards is the easiest and most predictable path to market.
Guidance = FDA expectation (usually! We’ll dive into this later)
99% of the time, conformance = faster overall time to market

## Slide 10: Experiences from the Field - Overview
Systems Engineering Background
Regulatory and Quality Role
  Design Assurance (“Systems Engineering Lite”)
  Remediated QMS and Business Processes
  Led FDA Inspections and Audit Prep (Front and Back Room)

## Slide 11: Pre-Market – Case Study 1
Major international medical device manufacturer
Acquired a “software shop” with non-medical and (surprise to them) medical software
No QMS to start
Acquired company was very adverse to anything outside of “pure agile”
Acquiring company was not knowledgeable on agile practices
Active med device development; 510(k) deadline already set in stone
Development Take-Aways:
  QMS, Training, Managed Expectations, Skirting the Grey Areas
FDA Review Take-Aways
  510(k) went through with minimal Ais
  Questions focused on risk, guidance document key points
  Deficiencies in DHF, DMR, and QMS remained.

## Slide 12: Pre-Market – Case Study 2
Major Med Device Manufacturer / Distributor
Innovation “spin off” – desired faster time to market
95% of hires have no medical device experience
Parent company had no medical device software experience; QMS focused on low-risk disposables
“I don’t want to document, I got hired on to code.”
Many standards and guidance that needed integrated to development already underway
Development Take-Aways:
  Quality plan to address lack of control
  Regulatory debt / regulatory stakeholders
  SE and SwE leadership with eventual transition to team
  Suggestion of new hires

## Slide 13: Post-Market – Case Study 1
Market leader in Class II in vivo blood handling products
Recently adopted agile (<2 years)
Under warning letters
Major DHF / DMR deficiencies, including software
Front Room Inspection Summary:
  Inspector – long career with military, transition to FDA
  “I don’t understand software at all”; enlisted help when needed from local office
  Looked towards FDA guidance on content of software in premarket submissions
  Was subject to the mercy of the inspected company WRT software

## Slide 14: Post-Market – Case Study 2
Same company as “pre-market – case study 1”
Product cleared a year earlier had its DHF & DMR inspected by FDA
Inspector competent in software and complex hardware systems
Front Room Inspection Summary:
  Items ignored during design (to get 510(k) cleared earlier) resulted in two 483s
  One 483 was related to the use of unvalidated tools to manage agile processes (JIRA)
    Verification referenced items in JIRA which should have been considered uncontrolled
  Inspector was familiar with scrum/agile, but only focused on outputs described in FDA guidance
  Inspector did not dive into control required for 2-week sprint configuration management

## Slide 15: The new Six Sigma?
The original authors of the Agile Manifesto have had much to say about the recent explosion of agile methodologies.
  Dave Thomas: The word “agile” has been subverted to the point where it is effectively meaningless, and what passes for an agile community seems to be largely an arena for consultants and vendors to hawk services and products.
  Andrew Hunt: [Agile] has become sloganized; meaningless at best, jingoist at worst. We have large swaths of people doing 'flaccid agile,' a half-hearted attempt at following a few select software development practices, poorly
Dave Thomas: I think it is time to retire the word “Agile.” […] Let’s develop with Agility.

## Slide 16: Our Strategy
Avoid the dogma of tools and practices
Understand the plight of engineers moving from unregulated to regulated environments
Let’s be as agile as practicable:
  within the constraints of regulatory bodies
  with the safety of the patients and clinicians as the cornerstone.
Illustrate “pay me now vs. pay me later” options:
  Engineer job security and specialization
  Cost of audits, inspections, 483s, etc.
  Unvalidated tools = adulterated product
  Can’t avoid documentation – even if someone else writes it
Stay within the guard rails!

## Slide 17: Summary
Auditors, inspectors, and reviewers are like snowflakes
Higher level problems invite subsequent levels of scrutiny
Consider pushing down agile practices into lower level QMS elements or guidance
510(k) clearance does not = a compliant QMS, DHF, or DMR!
Guard rails in place before design begins

## Slide 18: Questions and Discussion
Questions
Pre-Market Experiences
  Team Dynamics
  Submission Experiences
Post-Market Experiences
  Post-Market Inspections
  Maintenance of Product
  Updated presentation at cmdmedtech.com/incose